PENINSULA LABORATORIES, LLC

### A Member of the bachem group

GENERAL PROTOCOL

FOR

Peptide Enzyme Immunoassay (EIA)

CAUTION:  INVESTIGATIONAL DEVICE

LIMITED BY FEDERAL LAW TO INVESTIGATIONAL USE

FOR RESEARCH USE ONLY

NOT FOR USE IN DIAGNOSTIC PROCEDURES

EIA KITS - BASIC NOTIONS AND FACTS

Our EIA kits are competitive immunoassays whose working principle can be diagrammed as follows:

The antiserum is immobilized on a 96-well plate.  A constant concentration of Bt-tracer (biotynilated tracer) and varying concentrations of unlabeled standard or sample peptide compete for binding specifically to the antiserum. Captured Bt-tracer is subsequently bound by SA-HRP (streptavidin-conjugated horseradish peroxidase), which produces a soluble colored product after TMB (3,3’,5,5’-tetramethyl benzidine dihydrochloride) is added.

The sequence of the standard peptide is shown on the datasheet.  It is also in our catalog and on our web site (www.BACHEM.com) as the "antigen sequence"  (note that large protein sequences are usually not shown).

The standard is used to make a standard curve in the range specified in the kit's datasheet. Standard curves are S-shaped (on a semi-log plot) but for a few kits they appear to be almost linear over the kit's range. The measuring range is the range of standard concentrations near the middle or near the IC50 of the standard curve.  Typical standard curves for most kits are printed in our catalog but can also be obtained via e-mail (immunology@usbachem.com).  Unknown sample concentrations are measured by comparing its absorbance with the standard curve.

We include sufficient reagents for 96 determinations, enabling the generation of duplicate six- or ten-point standard curves and the analysis of up to 41 samples in duplicate.

VARIATION ( ACCURACY ( EXTRACTION ( CROSSREACTIVITY

The kit's IC50, or the shape of the standard curve, may show some variation but this will not affect the kit's accuracy in the measuring range.  The kit accurately measures sample peptides if the following conditions are met.

A) The kit's antiserum must not cross-react appreciably with other factors present in the sample. Cross-reactivity tables are included with each kit.  The user may wish to test the cross-reactivity with other peptides, some of which may be available from our catalog.

B) The sample peptides must be identical to the kit's standard.  Ideally the kit's synthetic standard mimics the natural peptide perfectly.  Sometimes, however, natural peptides exist as families of species related by a common or similar sequence.  In addition, natural peptides may be enzymatically or spontaneously modified, may exist in complexes, and may assume alternative structural forms. In these cases the kit might not measure the exact concentration of a particular natural peptide species, but, it may still be used for relative average measurements.

C) Serum samples must be extracted.  Factors present in serum can bind to kit components and affect the measurements.  Thus, it is advisable to extract peptides from serum before measuring them. Extraction may not be as important if the sample is diluted before measuring or if it is in tissue culture medium, but, in general, one may expect low sensitivity or high background problems if extraction is not used.  Some antigens can be measured directly in un-extracted serum.  If so, a specially formulated extraction-free kit may be available in our catalog or as a custom product, which can be used without extraction for human or rat serum.

D)  Both samples and standards must be measured under the same conditions.

TECHNICAL HELP and ORDERING INFORMATION

Please e-mail technical questions to immunology@usbachem.com

For additional kits or listings of our most current catalog products, please visit our web site at www.bachem.com or call us at:
Tel:		(800) 922-1516 (for U.S. and Canada)
		(650) 592-5392 (for International)
Fax:  	(650) 595-4071

SAFETY PRECAUTIONS

The physical and chemical properties of the reagents contained in this kit have been tested individually.  Reagents do not contain ingredients which have been determined to be health hazards and which comprise greater than 1% of the mixture or which could be released from the mixture in concentrations that would exceed OSHA permissible exposure limits.

Hazardous Ingredients

The standard is supplied in lyophilized form.  The lyophilized antiserum contains Tween-20 and thimerosal.  The EIA buffer concentrate contains Tris and thimerosal.  The buffer is in liquid form.  The SA-HRP contains 0.01% methylisothiazolone, 0.01% bromonitrodioxane, and 10 ppm Proclin 300 as a preservative.

Physical and Chemical Data

Components are stable in closed containers under normal temperatures and pressures.  No hazardous polymerization is known.

Fire and Explosion Data

Components are non-combustible with negligible fire hazard when exposed to heat or flame.  Fire fighting media should be appropriate to burning material.

Health Hazards

Components may be harmful by inhalation, ingestion, or skin absorption and may cause skin irritation or eye irritation.  In case of eye contact, flush eye with water and contact a physician.  In case of skin contact, wash skin with soap and water.

Reactivity Data

Components are stable in closed containers under normal temperatures and pressures.

Spill and Disposal Procedures

For spills, ventilate area and wash spill site.  For disposal, please dispose in accordance with local regulations.

Handling and Storage Information

Safety glasses, gloves, and a full-length lab coat should be worn to prevent unnecessary contact.

The above information is believed to be correct but does not purport to be all-inclusive and shall be used only as a guide.  It is the user’s responsibility to determine the suitability of this information for the adoption of safety precautions as may be necessary.  Peninsula Laboratories, LLC shall not be held liable for any damage resulting from the handling or use of the above product.

GUARANTEE AND LIMITATION OF REMEDY

PENINSULA LABORATORIES, LLC makes no guarantee of any kind, expressed or implied, which extends beyond the description of the material in this kit, except that these materials and this kit will meet our specifications at the time of delivery.  The customer’s remedy and PENINSULA LABORATORIES, LLC’s sole liability hereunder are limited at PENINSULA LABORATORIES, LLC’s option to refund the purchase price or the replacement of material that does not meet our specifications.  By the acceptance of our products, the customer indemnifies and holds PENINSULA LABORATORIES, LLC harmless against, and assumes all liability for the consequences of its use or misuse by the customer, its employees or others.

STORAGE

When you receive the kit, store the lyophilized components at a constant -20 °C for up to one year from the kit's assembly date.  The remaining components should be stored in the refrigerator (2-4 °C) also up to 1 year.   Long term storage, improper storage conditions and large temperature fluctuation cycles may result in the accumulation of precipitates in the TMB solution and in the EIA buffer concentrate.  These precipitates should not affect the assay noticeably.  Nevertheless, if you observe such precipitates, we recommend that you avoid them by allowing them to sink to the bottom.

MATERIALS PROVIDED

- EIA buffer concentrate (50 ml 20 X concentrate).

- One 96-well immunoplate with acetate plate sealer.

- One vial of peptide antibody (lyophilized powder).

- One vial of peptide standard (1 (g lyophilized powder).

- One vial of biotinylated peptide (lyophilized powder).

- Streptavidin-HRP concentrate (100 (l 200 X concentrate).

- Substrate solution (11ml of TMB and H2O2).

- 2 N HCl (15 ml)

- One vial standard diluent (peptide-free rat or human serum) for extraction-free kits only (8 ml)

.

MATERIALS NOT PROVIDED

- SEP-COLUMN containing 200 mg of C18 Cat No. Y-1000

- Extraction kit Cat No S-5000 for 50 samples with buffers

- 96-well microtiter plate reader set up to measure 450 nm.

- 96-well plate washer and shaker (optional)

- Sterile deionized or USP water

- Curve fitting software (optional)

- Test tubes, pipettes and various other standard laboratory items

ASSAY PROTOCOL

Reagents, samples, and the plate should be brought to room temperature before use.  Shakers (optional) may help lower the experimental variation of duplicates (recommended at 60 rpm).

I - LAYOUT

Design the plate layout. The following layouts are suggested for either 6-point or 10-point standard curves. The standards and the samples are duplicated, and two blank wells are included for background determinations.  These layouts accommodate 37 and 41 duplicate samples, respectively.

Six-Point Standard Curve Layout

|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| A | Blk | Blk | U2 | U2 | U10 | U10 | U18 | U18 | U26 | U26 | U34 | U34 |
| B | S1 | S1 | U3 | U3 | U11 | U11 | U19 | U19 | U27 | U27 | U35 | U35 |
| C | S2 | S2 | U4 | U4 | U12 | U12 | U20 | U20 | U28 | U28 | U36 | U36 |
| D | S3 | S3 | U5 | U5 | U13 | U13 | U21 | U21 | U29 | U29 | U37 | U37 |
| E | S4 | S4 | U6 | U6 | U14 | U14 | U22 | U22 | U30 | U30 | U38 | U38 |
| F | S5 | S5 | U7 | U7 | U15 | U15 | U23 | U23 | U31 | U31 | U39 | U39 |
| G | S0 | S0 | U8 | U8 | U16 | U16 | U24 | U24 | U32 | U32 | U40 | U40 |
| H | U1 | U1 | U9 | U9 | U17 | U17 | U25 | U25 | U33 | U33 | U41 | U41 |

Ten-Point Standard Curve Layout

|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| A | Blk | Blk | S8 | S8 | U6 | U6 | U14 | U14 | U22 | U22 | U30 | U30 |
| B | S1 | S1 | S9 | S9 | U7 | U7 | U15 | U15 | U23 | U23 | U31 | U31 |
| C | S2 | S2 | S10 | S10 | U8 | U8 | U16 | U16 | U24 | U24 | U32 | U32 |
| D | S3 | S3 | U1 | U1 | U9 | U9 | U17 | U17 | U25 | U25 | U33 | U33 |
| E | S4 | S4 | U2 | U2 | U10 | U10 | U18 | U18 | U26 | U26 | U34 | U34 |
| F | S5 | S5 | U3 | U3 | U11 | U11 | U19 | U19 | U27 | U27 | U35 | U35 |
| G | S6 | S6 | U4 | U4 | U12 | U12 | U20 | U20 | U28 | U28 | U36 | U36 |
| H | S7 | S7 | U5 | U5 | U13 | U13 | U21 | U21 | U29 | U29 | U37 | U37 |

Blk = blank   S = standards   U = unknown samples

II - PREPARE SAMPLES

1- Sample extraction.  Sample extraction is recommended especially for serum samples.  It may not be as important for some tissue culture samples.  See "Suggested Protocol for Sample Extraction" below for details.  The kit may still be used without extraction but this may cause unexpected results due to the possible binding between serum proteins and kit components.

If you purchased an Extraction-Free Kit you may use it for the measurement of human or rat serum or plasma (according to its designation) without performing an extraction.

2 - Sample concentration.  The concentration of the target molecule must be within the measuring range of the kit (in a region around the IC50).  If you cannot estimate the concentration range of your sample you can prepare it at different concentrations such that one of the samples may be within the measuring range.

III - PREPARE KIT COMPONENTS

Lyophilized kit components should not be re-hydrated until they are needed.  In particular, the standard and the tracer may not be needed until after an incubation step.  Please check the appropriate protocol.

1- Equilibrate unopened kit components to room temperature.  Do not open reagents and immunoplate while they are cold.  Equilibrate at room temperature before opening to avoid accumulation of moisture.

2 - Prepare the EIA buffer.  Dilute the EIA buffer concentrate to 1,000 ml with sterile deionized or USP pure water (18 MOhm) and mix well.

3 - Reconstitute the standard.  In protocol V the standard is used after an 1-hour incubation with antiserum .  Add 1 ml of standard diluent buffer to the vial of lyophilized standard peptide (1 (g)  and vortex.  If samples are to be extracted and re-suspended in EIA buffer as decribed below, use EIA buffer as a diluent.  Otherwise, we encourage customers to use their own diluent such that standards and samples will be treated equally.  Regardless, the "zero standard" should yield an OD close to that of the blank. For extraction-free kits we provide peptide-free rat or human serum as a diluent,  but again, customers should use their own sample diluent provided it does not bind appreciably to the antiserum.

4 - Reconstitute the antiserum.  Add 5 ml of EIA buffer and vortex.

5 - Reconstitute the Bt-tracer.  Please check the appropriate  protocol and reconstitute just it is needed.  Add 5 ml of EIA buffer to the vial of lyophilized biotinylated peptide and vortex.

6 - Prepare the standards. Make serial dilutions of the standard to cover the range of this kit.  The following tables may be useful to help you prepare either a 6-point or a 10-point standard curve (more precision).

IV - ASSAY

There are six different protocols: I, II, III, IV, V, and VI. Please check the datasheet enclosed with the kit and choose the appropriate protocol.

PROTOCOL I (Std.Ab.Bt)

1 - Into each well of the immunoplate add

50 (l standard or sample

25 (l antiserum

25 (l Bt-tracer

Add only buffer to the blank wells (these will give you the background to be subtracted from all readings).  These should not be confused with the "S0 Standards" which contain no standard peptide and which will yield the highest OD readings.

2 - Incubate at room temperature for 2 hours.

3 - Wash immunoplate 5 times with 300 (l/well of EIA buffer.  Be very careful not to cross-contaminate between wells in the first wash/dispensing cycle.  In each wash cycle empty plate contents with a rapid flicking motion of the wrist, then gently blot dry the top of plate on paper towels.   Dispense 300  (l of EIA buffer into each well and gently shake for at least a few seconds.  Thorough washing is essential.

4 - Add 100 (l /well of streptavidin-HRP.  Tap or centrifuge the SA-HRP vial to collect all liquid contents on the bottom of the vial.  Dilute 1/200 in EIA buffer (60 (l / 12 ml) and vortex. Add 100(l to all wells, including the blanks.

5 - Incubate at room temperature for 1 hour.

6 - Wash immunoplate 5 times (see step 3).

7 - Add 100 (l/well of TMB solution. Add to all wells, including the blanks.

8 - Incubate at room temperature (usually 30 - 60 minutes).  You may read the developing blue color at 650 nm and use the data for your calculations

9 - Terminate reactions by adding 100 (l 2 N HCl per well .

10 - Read absorbance at 450 nm within ten minutes.

PROTOCOL II (Std.Ab1hr.BtON)

1 - Into each well of the immunoplate add

50 (l standard or sample

25 (l antiserum

Add only buffer to the blank wells (these will give you the background to be subtracted from all readings).  These should not be confused with the "S0 Standards" which contain no standard peptide and which will yield the highest OD readings.

2 - Incubate at room temperature for 1 hour.

3 - Rehydrate the Bt-tracer (see III-5) and add 25 (l / well

Add Bt-tracer to all wells except the blank wells.

4 - Incubate at 4(C overnight.

Shorter incubation times may result in low signal.

5 - Wash immunoplate 5 times with 300 (l/well of EIA buffer.  Be very careful not to cross-contaminate between wells in the first wash/dispensing cycle.  In each wash cycle empty plate contents with a rapid flicking motion of the wrist, then gently blot dry the top of plate on paper towels.   Dispense 300  (l of EIA buffer into each well and gently shake for at least a few seconds.  Thorough washing is essential.

6 - Add 100 (l /well of streptavidin-HRP.  Tap or centrifuge the SA-HRP vial to collect all liquid contents on the bottom of the vial.  Dilute 1/200 in EIA buffer (60 (l / 12 ml) and vortex. Add 100(l to all wells, including the blanks.

7 - Incubate at room temperature for 1 hour.

8 - Wash immunoplate 5 times (see step 5).

9 - Add 100 (l/well of TMB solution. Add to all wells, including the blanks.

10 - Incubate at room temperature (usually 30 - 60 minutes).  You may read the developing blue color at 650 nm and use the data for your calculations

11 - Terminate reactions by adding 100 (l 2 N HCl per well .

11 - Read absorbance at 450 nm within ten minutes.

PROTOCOL III (Std.Ab1hr.Bt)

1 - Into each well of the immunoplate add

50 (l standard or sample

25 (l antiserum

Add only buffer to the blank wells (these will give you the background to be subtracted from all readings).  These should not be confused with the "S0 Standards" which contain no standard peptide and which will yield the highest OD readings.

2 - Incubate at room temperature for 1 hour.

Shorter pre-incubations may result in lower sensitivity.

3 - Rehydrate the Bt-tracer (see III-5) and add 25 (l / well

Add Bt-tracer to all wells except the blank wells.

4 - Incubate at room temperature for 2 hours.

5 - Wash immunoplate 5 times with 300 (l/well of EIA buffer.  Be very careful not to cross-contaminate between wells in the first wash/dispensing cycle.  In each wash cycle empty plate contents with a rapid flicking motion of the wrist, then gently blot dry the top of plate on paper towels.   Dispense 300  (l of EIA buffer into each well and gently shake for at least a few seconds.  Thorough washing is essential.

6 - Add 100 (l /well of streptavidin-HRP.  Tap or centrifuge the SA-HRP vial to collect all liquid contents on the bottom of the vial.  Dilute 1/200 in EIA buffer (60 (l / 12 ml) and vortex. Add 100(l to all wells, including the blanks.

7 - Incubate at room temperature for 1 hour.

8 - Wash immunoplate 5 times (see step 5).

9 - Add 100 (l/well of TMB solution. Add to all wells, including the blanks.

10 - Incubate at room temperature (usually 30 - 60 minutes).  You may read the developing blue color at 650 nm and use the data for your calculations

11 - Terminate reactions by adding 100 (l 2 N HCl per well .

11 - Read absorbance at 450 nm within ten minutes.

PROTOCOL IV (Std.AbON.Bt)

1 - Into each well of the immunoplate add

50 (l standard or sample

25 (l antiserum

Add only buffer to the blank wells (these will give you the background to be subtracted from all readings).  These should not be confused with the "S0 Standards" which contain no standard peptide and which will yield the highest OD readings.

2 - Incubate at 4(C overnight.

Shorter pre-incubations may result in lower sensitivity.

3 - Incubate at RT for 1 hour.

This is to allow the plate to reach room temperature.

3 - Rehydrate the Bt-tracer (see III-5) and add 25 (l / well

Add Bt-tracer to all wells except the blank wells.

4 - Incubate at room temperature for 2 hours.

5 - Wash immunoplate 5 times with 300 (l/well of EIA buffer.  Be very careful not to cross-contaminate between wells in the first wash/dispensing cycle.  In each wash cycle empty plate contents with a rapid flicking motion of the wrist, then gently blot dry the top of plate on paper towels.   Dispense 300  (l of EIA buffer into each well and gently shake for at least a few seconds.  Thorough washing is essential.

6 - Add 100 (l /well of streptavidin-HRP.  Tap or centrifuge the SA-HRP vial to collect all liquid contents on the bottom of the vial.  Dilute 1/200 in EIA buffer (60 (l / 12 ml) and vortex. Add 100(l to all wells, including the blanks.

7 - Incubate at room temperature for 1 hour.

8 - Wash immunoplate 5 times (see step 5).

9 - Add 100 (l/well of TMB solution. Add to all wells, including the blanks.

10 - Incubate at room temperature (usually 30 - 60 minutes).  You may read the developing blue color at 650 nm and use the data for your calculations

11 - Terminate reactions by adding 100 (l 2 N HCl per well .

11 - Read absorbance at 450 nm within ten minutes.

PROTOCOL V (Ab1hr.Std2hr.BtON)

1 - Into each well of the immunoplate add 25 ul antiserum.

Add only buffer to the blank wells (these will give you the background to be subtracted from all readings).  These should not be confused with the "S0 Standards" which contain no standard peptide and which will yield the highest OD readings.

2 - Incubate at room temperature for 1 hour.

3 - Add 50 ul standard or sample.

Do not wash plate before adding.  Do not add to blank wells.

4 - Incubate at room temperature for 2 hours.

Shorter pre-incubations may result in lower sensitivity.

5 - Rehydrate the Bt-tracer (see III-5) and add 25 (l / well

Add Bt-tracer to all wells except the blank wells.  Agitate gently.

6 - Incubate at 4(C overnight.

7 - Wash immunoplate 5 times with 300 (l/well of EIA buffer.  Be very careful not to cross-contaminate between wells in the first wash/dispensing cycle.  In each wash cycle empty plate contents with a rapid flicking motion of the wrist, then gently blot dry the top of plate on paper towels.   Dispense 300  (l of EIA buffer into each well and gently shake for at least a few seconds.  Thorough washing is essential.

8 - Add 100 (l /well of streptavidin-HRP.  Tap or centrifuge the SA-HRP vial to collect all liquid contents on the bottom of the vial.  Dilute 1/200 in EIA buffer (60 (l / 12 ml) and vortex. Add 100(l to all wells, including the blanks.

9 - Incubate at room temperature for 1 hour.

10 - Wash immunoplate 5 times (see step 5).

11 - Add 100 (l/well of TMB solution. Add to all wells, including the blanks.

12 - Incubate at room temperature (usually 30 - 60 minutes).  You may read the developing blue color at 650 nm and use the data for your calculations

13 - Terminate reactions by adding 100 (l 2 N HCl per well .

14 - Read absorbance at 450 nm within ten minutes.

PROTOCOL VI (S.Ab.BtON)

1 - Into each well of the immunoplate add

50 (l standard or sample

25 (l antiserum

25 (l Bt-tracer

Add only buffer to the blank wells (these will give you the background to be subtracted from all readings).  These should not be confused with the "S0 Standards" which contain no standard peptide and which will yield the highest OD readings.

2 - Incubate at 4(C overnight.

3 - Wash immunoplate 5 times with 300 (l/well of EIA buffer.  Be very careful not to cross-contaminate between wells in the first wash/dispensing cycle.  In each wash cycle empty plate contents with a rapid flicking motion of the wrist, then gently blot dry the top of plate on paper towels.   Dispense 300  (l of EIA buffer into each well and gently shake for at least a few seconds.  Thorough washing is essential.

4 - Add 100 (l /well of streptavidin-HRP.  Tap or centrifuge the SA-HRP vial to collect all liquid contents on the bottom of the vial.  Dilute 1/200 in EIA buffer (60 (l / 12 ml) and vortex. Add 100(l to all wells, including the blanks.

5 - Incubate at room temperature for 1 hour.

6 - Wash immunoplate 5 times (see step 3).

7 - Add 100 (l/well of TMB solution. Add to all wells, including the blanks.

8 - Incubate at room temperature (usually 30 - 60 minutes).  You may read the developing blue color at 650 nm and use the data for your calculations

9 - Terminate reactions by adding 100 (l 2 N HCl per well .

10 - Read absorbance at 450 nm within ten minutes.

V - ANALYSIS

Plot data and calculate results. We recommend that you use curve fitting software for your data analysis.  Plate readers often include such software packages.  This is, however, not essential and you may opt to plot manually on semi-log paper.  You can also use a spreadsheet program.  Should you need help with the latter method we recommend the following procedure.

Set up a spreadsheet as shown below (note that the values on the spreadsheet are merely illustrative and are not necessarily typical for this particular kit).  If you e-mail us at immunology@usbachem.com we will be happy to send you the actual working Excel spreadsheet shown below.

Set up an 8 X 12 area to match the layout of the plate (in this case it is the six-point layout) and dump the plate reader data in it.

Calculate the average of the blanks on another cell as indicated by the arrows starting from cells A1 and A2.

Enter the concentration of the standards (see under ng/ml in figure).

Calculate the average of the ODs of the standards and subtract the background (blank) as indicated by the arrows for the last standard (cells G1 and G2).

Plot a standard curve points on a semi-log scale.   Use the y axis for the average of the OD readings (minus bkg) and the x axis for the standard concentrations in ng/ml.

Use the equation shown below to calculate the values on the "FIT" column and plot a smooth line of FIT values versus standard concentrations.  Then change the parameters a (max), b (slope), c (IC50), and d (min), until you are satisfied that fit is good.

Next calculate the average of your sample readings and subtract the blank average (see arrows starting from A3 and A4, and the arrows leading to "Average").

Finally, you may use the "reverse" of the equation above to calculate the concentrations in ng/ml for all your samples.

Caution: In the example above we set the S0 ng/ml value at 0.001 so that the value could be plotted on a semi-log scale, but value is actually ZERO.  Furthermore, when you calculate sample concentrations using the "reverse" equation if y = d or  y > a or y < d, the reading is out of range and the calculation will yield an error or a meaningless negative concentration.

SUGGESTED PROTOCOL FOR SAMPLE EXTRACTION

We have provided an excess amount of standard that you may use to determine if extraction is required.  For example, if you are working with serum, you may spike it with known amounts of standard and check if they are accurately determined by the assay with and without extraction.  Extraction eliminates potentially interfering substances, such as albumin.  Extraction may also be necessary to concentrate the sample to within the measuring range. As with any purification technique, recovery of the desired substance is likely to be incomplete.  Therefore, both optimization and quantification of the extraction procedure are recommended for more accurate determinations.  While we cannot provide you with extraction optimization and quantification protocols, we have included enough standard in the kit should you wish to use it for this purpose.

C18 Sep-Column Extraction Method

The following generic protocol is meant to help users with little experience in extracting their samples.  It should be applicable to different biological fluids but should not be thought of as an optimized protocol for any particular antigen.

Required Materials

- SEP-COLUMN containing 200 mg of C18 Cat No. Y-1000

- Buffer A (BUFF-A): 1% trifluoroacetic acid (TFA, HPLC Grade). (Acidifies plasma sample to remove interfering proteins such as albumin)

- Buffer B (BUFF-B): 60% acetonitrile (HPLC Grade), 1% TFA, and 39% distilled water. (Elutes peptide from C18 column)

- You may also consider purchasing Extraction kits (Cat No. S-5000), which include SEP-columns and buffers

Withdrawal and Preparation of Plasma

- Collect blood samples (2 - 6 ml) into a chilled syringe and transfer into a polypropylene tube containing EDTA  (1 mg/ml of blood) as an anticoagulant and Apronitin (500 KIU/ml of blood) as a protease inhibitor  at 4°C.  Do not use heparinized tubes as they may interfere with the assay.  Vacutainers with EDTA are acceptable.

- Centrifuge blood at 1,600xg for 15 minutes at 4°C.

- Collect the top (plasma) layer.

- Proceed to extraction immediately or freeze at -70°C for later use.

Extraction Procedure

- Add an equal amount of Buffer A to the plasma.

- Centrifuge at 6,000xg to 17,000xg for 20 minutes at 4°C.

- Transfer supernatant to a new tube discarding any pellet that may be present.

- Equilibrate a SEP-COLUMN by washing with 1 ml Buffer B followed by 3  X  3 ml Buffer A.

- Load the plasma solution onto the equilibrated SEP-Column.

- Slowly wash the column with Buffer A (3 ml, twice) and discard the wash. A light vacuum (10 sec/drop) may be applied to the column.

- Elute the peptide slowly with Buffer B (3 ml, once) and collect eluant in a polypropylene tube.  A light vacuum may be applied as in previous step.

- Freeze-dry eluant to dryness using a dry ice/methanol bath to freeze the sample and a centrifugal concentrator to evaporate it

- Dissolve the residue in a suitable volume of EIA buffer such that the concentration of the substance of interest will fall close to the IC50 (within the measuring range).

TROUBLESHOOTING

Most problems are caused by poor technique or by alterations of the protocol.   Also, please check that the kit's expiration date has not passed and that the kit has been stored properly.  Please read the storage section on page 5.

- Can the kit be used more than once?

  - Although we do not guarantee the performance of our kits on a second or subsequent use, the end user should be able to use the kit multiple times if the reconstituted specific anti-serum, standard and Bt-tracer are stored at or below a constant -20 °C and the rest of the components are kept dry and refrigerated (2-4 °C).  Freezing aliquots of the reconstituted components may further extend multiple use lifetime.

- What are some sources of inaccurate readings?

  - exceeding the OD range of plate reader

  - dirt or grease on the bottom of the plate - Wipe with 70% ethanol

  - air bubbles or foaming in wells.

- The standard curve does not look right.

  - If you wait too long to read, the curve will be flattened at the top. If you are not familiar with the kit we recommend you read the plate several times while the blue color is still developing (at 650 nm).   The blue color will be less intense compared to that of the terminated reactions (Yellow 450 nm) but the data are still good and this way you won't risk losing the lower end of the range.

  - Some curves are almost linear, check our catalog for a typical plot of the product you are using - it may be normal for that particular product.

- The IC50 is not as expected.

  - A difference by a factor of two or three may be normal for some kits and may be caused by the time it takes to equilibrate the binding of the tracer and the standard.  This will be especially true for pre-incubation protocols.  Otherwise, it is possible that the standard degraded.  If it is possible you should always include your own reliable standard at a concentration close to the expected IC50 to check the accuracy of the kit.

  - In cases where the standard curves are almost rectilinear, accurate IC50 values cannot be calculated.

  - Using excessive amounts of antiserum or tracer, or using a degraded standard may all raise the IC50.

- There is too much variation in duplicated readings.

  - There are only trivial explanations for this such as: (a) poor mixing, (b) poor pipetting technique or faulty pipettes, (c) kit reagents not allowed to equilibrate to room temperature before use, (d) cross-contamination of samples, e.g. droplets or spray from one well to the next, (e) bubbles or foaming in the wells, or finger prints or dirt on the bottom of the plate, etc.

- The ODs are lower than expected.

  - The color intensity has little to do with the accuracy of the kit, as long as the slope in the measuring range is normal, but, if the intensity is extremely low, and assuming that you have waited long enough, this may mean that one of the components (antiserum, Bt-tracer, SA-HRP, or TMB) was added in low amounts or was degraded due to incorrect storage or excessive freeze-thawing.

- The curve looks OK but the results seem implausible.

  - Possibly you used different solvents or conditions for standards and samples.

  - The antiserum may bind to another antigenically similar peptide

  - Antigen was lost during extraction or extraction did not eliminate interfering factors

  - Make sure that the kit's antigen is the same as the target that you are trying to measure.  Sometimes the kit's antigen is a peptide that is part of but not the complete natural protein.   If so the kit can still be used for determining relative concentrations but not necessarily for determining the absolute concentration of the complete protein antigen.

REFERENCES

T. Porstmann and S.T. Kiessig Enzyme immunoassay techniques.  An Overview. J. Immunol. Methods 150, 5-21 (1992)

S. Avrameas Amplification systems in immunoenzymatic techniques.

J. Immunol. Methods 150, 23-32 (1992)

E. Bucht et al. A rapid extraction method for serum calcitonin.

Clin. Chim. Acta, 195, 115-124 (1991)

NOTES

NOTES

# Text boxes

SA-HRP

TMB

Bt-Tracer

Stopped  Yellow

450 nm  Color

2 N HCL

Developing  Blue 650 nm Color

Specific

Antiserum

Coated

Plate

Standard

or Sample Peptide